## Slide 1: V-Disk Automated Extracellular Vesicle Isolation
Dr. Jan Lüddecke
03.11.2025
ELBS general assembly
Madrid

## Slide 2: Hahn-Schickard – R&D in microsystems engineering
Microfluidic Platforms
Bioanalytics
Laboratory Automation
Notes: https://www.youtube.com/watch?v=rYVTH61jNW8

## Slide 3: State of the art EV purification
Ultracentrifugation
SEC
Precipitation
Time consuming, many manual steps, lots of different protocols
Single methods usually not sufficient
+SEC
+UF
+TFF +UF
+DGUC
+UF
EV purification needs to be standardized and automated for clinical application
Whole blood  purified EVs

## Slide 4: EVs – Needle in a haystack
Zarovni et al., J. Extracell. Vesicles, 2025
Notes: According to the literature, on average, there are about 10¹⁶ particles in 1 mL of plasma.

## Slide 5: Label free EV purification
Negative
Positive
0
10
20
40
60
80
500
Diameter (nm)
Surface charge
LDL
VLDL
HDL
EV
Size based purification
Charge based purification
Adapted from Van Deun et al., 2020
Protein
Notes: On the other hand, most label-free EV isolation techniques rely on either charge or size. However, as you can see, the mentioned contaminants overlap with EVs in at least one of these properties, that is why it’s often necessary to combine both to achieve a high-purity isolation. For example, if a method separates particles first by charge and then by size, the final isolate will be more or less free of contaminants. 40 seconds

## Slide 6: Multi-modal EV isolation protocol
EV
MMC resin
Cation-exchange Chromatography
(CEX)
1 µm filtration
20 nm filtration
Mixed-mode Chromatography
Charge
Size
C+S
  Removal of (+) charged particles
  Enhances filtration
  Removal of particles >1 µm, <20 nm
  EV concentration
  Removal of remaining small particles
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+
EV
-
+
+
LDL
VLDL
+
EV
HDL
Proteins
Platelets
Proteins
CEX
CEX

## Slide 7: V-Disk + Disk-Player
+
Centrifugal microfluidics allows automation of complex biochemical assays

## Slide 8: V-Disk
  Disposable cartridge
  0.1 – 1 ml Plasma  700 µl purified EVs
  Thermal integration of filter membranes
  Dried storage chromatography resins
Filter-1 1 µm
Filter-2 20 nm
Dry resin
Reconstituted column
Notes: Add scale

## Slide 9: V-Disk – EV Morphology
  Cryo-EM analysis
  EVs are intact
V-Disk
SEC

## Slide 10: V-Disk vs SEC - Recovery
  GFP-EVs spiked into plasma
  80 % recovery
Isolate
Waste
SEC
V-Disk

## Slide 11: Reproducibility and Robustness
Created with BioRender.com
VDisk
SEC
Reproducibility assessment
Robustness assessment
Fasting
Healthy donor (n=3)
Postprandial
Healthy donor (n=3)
Lipemic
Non-lipemic
Notes: To further investigate the reproducibility of the VDisk, we tested it using three different biological replicates from healthy donors. Additionally, to evaluate the robustness of our platform and its ability to handle lipemic samples, we collected blood from the same donors on the same day, a few hours after they had consumed fatty food.

## Slide 12: EV size distribution – NTA
V-Disk
SEC
Fasting
Postprandial
  Particles in expected size range
  Size distribution unaffected by lipemic sample
Notes: Also the particle size distribution remained unchanged in VDisk isolates between fasting and postprandial plasma

## Slide 13: Contaminant removal
VDisk
SEC
Total protein
High-density lipoprotein (HDL)
Low- and very-low-density lipoprotein (V/LDL)
Fasting
Postprandial
All methods removed contaminants with similar efficiency in both fasting and postprandial plasma.
VDisk
SEC
VDisk
SEC
Protein removal %
ApoA1 removal %
ApoB100 removal %
Notes: However, all three methods methods removed contaminants with similar efficiency in both fasting and postprandial plasma

## Slide 14: Reproducibility and Robustness
CV≤15
CV≥30
15≤CV≤30
EV markers
Lipoprotein markers
Notes: We analyzed the EV isolates from each method using various biomarkers, including CD81, CD9, total protein, total ApoA, and total ApoB. Both VDisk configurations showed CV values below 15%, while SEC showed higher variability, reaching up to 30% in some cases.

## Slide 15: Impact of sample volume on V-Disk efficiency
1 mL
0.5 mL
0.1 mL
Inlet plasma sample volume
VDisk
EV isolate 700 µL
Notes: Until now, we have always evaluated the VDisk using 1 mL of plasma. However, we thought that would be interesting to see how changing the initial plasma volume might affect the VDisk’s performance. Therefore, we tested three different input volumes — 1 mL, 0.5 mL, and 0.1 mL of plasma

## Slide 16: Impact of sample volume on V-Disk efficiency
Yield
(EV/ sample vol.)
Purity
(total EV/ total protein)
Purity
(total EV/ total HDL)
1.0 mL
0.5 mL
0.1 mL
SEC
VDisk
1.0 mL
0.5 mL
0.1 mL
SEC
VDisk
1.0 mL
0.5 mL
0.1 mL
SEC
VDisk
Notes: Mention CD9 also gave us similar results
Signal to contaminant ratio
Membrane overloaded

## Slide 17: Impact of sample volume on V-Disk Efficiency
1.0 mL
0.5 mL
0.1 mL
1.0 mL
0.5 mL
0.1 mL
SEC
VDisk
Purity
(total EV/ (V/LDL)
Filtration time
Notes: Mention the result from other CD signals were similar

## Slide 18: Outlook: V-Disk v3
  Sample: up to 3 ml of whole blood or 1 ml of plasma
  Plasma generation and platelet removal
  Improved workflow
  Detection module for QC or EV markers

## Slide 19: Summary
  Automated EV purification from blood to concentrated EVs
  Robust, high recovery, high purity
  We are looking for:
  Feedback, user specification
  Project partners, interested in testing
  Commercialization partners

## Slide 20
Acknowledgments
Gustav Grether
Ehsan Arjmand
Nils Paust
Andreas Baur
Jan van Deun
Daniel Kainz
Prototyping @Hahn-Schickard
Precision milling
Hot embossing
Liquid Biopsy @Hahn-Schickard
Jan.Lueddecke
@Hahn-Schickard.de
Irina Nazarenko
John Atanga
Gonzalo Bustos
Notes: That brings me to the end of my presentation.
I’d like to thank my great team at Hahn-Schickard , as well as our project partners at the University Clinics in Freiburg and Erlangen. Also a big thanks to our prototyping team for their fantastic help.
If you’re interested in our work on liquid biopsy and extracellular vesicles, please check out our homepage.
And at the end thanks a lot for your attention.